## Slide 1: Pacemaker Timing Cycles
Dr Jassim kp

## Slide 2: References
Cardiac Pacing and ICDs by Kenneth A. Ellenbogen and Karoly Kaszala , seventh edition
Braunwalds heart disease- twelfth edition
Clinical Cardiac Pacing, Defibrillation and Resynchronization Therapy (Fifth Edition)
Barold, S & Garrigue, Stéphane & Clémenty, Jacques. (2002). Far-Field P-Wave Sensing by the Right Ventricular Lead of Conventional Dual Chamber Pacemakers. Journal of interventional cardiac electrophysiology : an international journal of arrhythmias and pacing. 6. 77-80. 10.1023/A:1014132408411. ,
.

## Slide 3
Mulpuru SK, Madhavan M, McLeod CJ, Cha YM, Friedman PA. Cardiac Pacemakers: Function, Troubleshooting, and Management: Part 1 of a 2-Part Series. J Am Coll Cardiol. 2017 Jan
Liaquat MT, Ahmed I, Alzahrani T. Pacemaker Malfunction. 2023 Jul 30. StatPearls Publishing; 2024 Jan–. PMID
Richter S, Muessigbrodt A, Salmas J, Doering M, Wetzel U, Arya A, Hindricks G, Brugada P, Israel CW. Ventriculoatrial conduction and related pacemaker-mediated arrhythmias in patients implanted for atrioventricular block: an old problem revisited. Int J Cardiol. 2013 Oct 09;168(4):3300-8
Martow E, Klein A, Green MS. Sense and sensitivity: P-wave oversensing causing inhibition in a dual-chamber pacemaker and practical considerations. HeartRhythm Case Rep. 2021 May 15

## Slide 4: Objectives
Describe the NBG Code
List the possible pacing modes
Explain timing intervals in single and dual chamber devices
Define blanking and refractory periods
Describe upper rate behavior
Pacemaker mediated tachycardia

## Slide 5
All potential variations of a single complete pacing cycle
Each timing circuit
Completes its cycle
Release of a pacing stimulus
Initiation of another timing cycle
Given timer resets
Starts the timing period again
Notes: the time from paced ventricular beat to paced ventricular beat; the time from paced ventricular beat to an intrinsic ventricular beat, whether it is a conducted R wave or a premature ventricular contraction (PVC); the time from paced atrial beat to paced atrial beat; the time from intrinsic atrial beat to paced atrial beat; the time from intrinsic ventricular beat to paced ventricular beat; and so forth
These cycles include events sensed, events paced, and periods when the sensing circuit or circuits are refractory

## Slide 6: PACING NOMENCLATURE - NBG Codes
I | II | III | IV | v
Chamber(s) Paced | Chamber(s) Sensed | Response to Sensing | Rate Modulation | MULTISITE PACING
O = None A = Atrium V = Ventricle D = Dual (A + V) S = Single (A or V) | O = None A = Atrium V = Ventricle D = Dual (A + V) S = Single (A or V) | O = None T = Triggered I = Inhibited D = Dual (T + I) | O = None R = Rate modulation | O = None A = Atrium V = Ventricle D = Dual (A + V)
Notes: A three-letter code describing the basic function of the various pacing systems was first proposed in 1974 by a combined task force from the American Heart Association and the American College of Cardiology.
NBG is a standard created by the NASPE and the British Pacing Group Modes
The third position refers to the mode of sensing, or how the pacemaker responds to a sensed event
. An I indicates that a sensed event inhibits the output pulse and causes the pacemaker to recycle for one or more timing cycles. T means that an output pulse is triggered in response to a sensed event
An R in the fourth position indicates that the pacemaker incorporates a sensor to control the rate independently of intrinsic electrical activity of the heart. The fifth position indicates whether multisite pacing is not present (O) or present in the atrium (A), ventricle (V), or both (D). Multisite pacing is defined for this purpose as stimulation sites in both atria, both ventricles, more than one stimulation site in any single chamber, or any combination of these.

## Slide 7: Rate Response / rate-adaptive pacing
Pacing rate response to physical, mental, emotional exertion
Improves exercise capacity in patients with chronotropic incompetence
Sensors- Accelerometer (mc)  physical motion
Physiological sensors minute ventilation, cardiac contractility, blood temperature, Ph, or volume changes
Notes: Increase in the pacing rate
The most common sensor used to drive rate response is the accelerometer, which rapidly detects physical motion.
Although more specific, they introduce greater latency and are often used with accelerometers to provide a blended response

## Slide 8
Limitations-
Accelerometer
need for upper extremity motion to drive the heart rate
Abrupt deceleration after a period of moderate-intensity exercise when the metabolic demands are still high
Notes: (so that walking with limited arm motion on a treadmill leads to a poor rate response, patients with tremor or car travelling in a bumpy roads)

## Slide 9
Minute ventilation sensors
variations in pulmonary impedance -> transmit a low-energy current
Not suitable -tachypnea related to pulmonary disease
Connected to ECG monitors- inappropriately increase the heart rate
Notes: Minute ventilation sensors -transmit a low-energy current from the lead tip to the pulse generator to measure the variations in pulmonary impedance
inappropriately increase the heart rate when patients are connected to ECG monitors that inject small currents into the body, which results in sensor error

## Slide 10: Primary functions
Pacing
Sensing

## Slide 11
Pacemaker timing cycle should be considered in milliseconds
60,000 msec = 1 min
60,000msec/ 60bpm = 1000msec

## Slide 12: Single chamber timing
Lower rate limit
Refractory period
Blanking period

## Slide 13: Single chamber timing
LRL- lower rate limit
Base rate or slowest rate that the pacemaker will allow the patient's heart to go.
programmable.
Restarts with each paced or sensed beat
Intrinsic HR > LRL : pacemaker will inhibit pacing and sense
Intrinsic HR < LRL : pacemaker will pace at LRL

## Slide 14: Refractory period
Period of time following a sensed or paced beat during which pacemaker will not respond to the sensed events.
Atrial refractory period-ARP ( AAI mode)
Ventricular refractory period- VRP
Notes: All pacing modes that can sense cardiac events must include blanking and refractory periods in their basic timing cycle
refractory after a paced or sensed ventricular event
event occurring within the VRP is not sensed and does not reset the ventricular timer

## Slide 15: Blanking period
Blanking period -period during which the sensing amplifier is “off” or “blind” to any cardiac event
Atrial blanking period (ABP)
Ventricular blanking period (VBP)
Notes: essential to the appropriate pacemaker function as they prevent sensing of known but clinically inappropriate signals, such as the evoked potential and repolarization
Once this period ends, the sense amplifier becomes alert and is receptive to the detection of native signals
Atrial blanking period (ABP) Atrial-sensing amplifier is “blind” and will not detect or respond to any atrial-sensed event Ventricular blanking period (VBP) Ventricular-sensing amplifier is “blind” and will not detect or respond to any ventricularsensed eve

## Slide 16: VVI
Ventricle - paced
Ventricle- sensed
sensed event- will inhibit the next pace
VVI 60 = Lower Rate timer of 1000 ms
    Pacing every 1 second if not inhibited
VP
VS
VP
VP
Lower rate timer 1000 ms
x
Lower rate timer 1000 ms
Notes: 16

## Slide 17: VVI
VRP | 
LRL | 
Escape intervarl
VBP

## Slide 18: Hysteresis
Promotes intrinsic activity below the programmed Lower Rate.
Prevents the device from overriding slow, but appropriate, intrinsic rhythms that may develop from extended periods of inactivity such as sleep
Only single chamber modes and with Rate Response Off

## Slide 19
a sensed event temporarily suspends the Lower Rate, and the pacemaker determines its escape rate from the hysteresis rate instead

## Slide 20: VOO - Asynchronous Pacing
Chamber paced - Ventricle
Chamber sensed - None
Response to sensing- None
1000 ms
1000 ms
1000 ms
VP
VP
VP
VP
Notes: 20
e simplest of all pacing modes
VOO results in fixed-rate pacing in the ventricle.
The intrinsic ventricular event cannot be sensed, and thus, does not interrupt the pacing interval.

## Slide 21: AAI Atrium - sensed Atrium- paced Response- inhibition
ARP | 
LRL | 
ABP

## Slide 22: AAI
Ventricular event not sensed
Exception – FAR FIELD SENSING
Notes: ventricular signal is large enough to be inappropriately sensed by the atrial lead, In this situation, the atrial timing cycle is reset
Sometimes this anomaly can be corrected either by making the atrial channel less sensitive or by lengthening the refractory period.

## Slide 23: Dual chamber timing
AV Interval
Pacing
  Tracking
Upper Tracking Rate (UTR)
Upper Sensor Rate (USR)
Far field
Crosstalk

## Slide 24
Basic dual chamber timing rests on 1 of 2 schemes
A-A or V-V timing
ventricular sensed or paced event atrial event = AEI, or VAI
( atrial escape interval , ventriculoatrial interval)
atrial sensed or paced event ventricular event= AV interval
atrial sensed event that occurs before completion of the AEI promptly terminates this interval and initiates an AVI, and the result is P-wave synchronous ventricular pacing.
intrinsic sinus rate < programmed LRL, AV sequential pacing at the programmed rate
Notes: interval from a ventricular sensed or paced event to an atrial paced event and is known as the AEI, or VAI
interval begins with an atrial sensed or paced event and extends to a ventricular event. This interval may be defined by a paced AV, PR, AR, or PV interval
atrial sensed event that occurs before completion of the AEI promptly terminates this interval and initiates an AVI, and the result is P-wave synchronous ventricular pacing.
8 If the intrinsic sinus rate is less than the programmed LRL, AV sequential pacing at the programmed rate or functional single-chamber atrial (AR) pacing occurs

## Slide 25: A-A Timing
Atrial rate = 60 ppm
Ventricular rate = 63 ppm (first interval) ; 57 ppm (second interval)
A-A is fixed and V-A auto adjusts
AV = 200
A-A = 1000 ms
A-A = 1000 ms
A-A
Timing
AV = 200
AV = 150
V-A = 850
V-A = 800

## Slide 26: V-V Timing
V-A is fixed and A-A auto adjusts
Atrial rate = 60; 63 ppm
Ventricular rate = 63; 60 ppm
A to A = 1000 ms
A to A = 950 ms
V-V
Timing
PAV = 200
V-A = 800
PR = 150
V-A = 800
PAV = 200

## Slide 27
Modified A-A - mainly use atrial based timing unless events such as PVCs occur, when the V-A interval is used to determine the timing of the next A pace.
Notes: is standard today where devices

## Slide 28: AV sequential pacing- ApVp
AV interval | 
VA interval | 
LRL | 
Ap
Vp
VBP
VRP
ABP
ARP
Notes: VRP prevents sensing of the evoked potential and the resultant T wave on the ventricular channel of the pacemaker

## Slide 29
A sensed or paced ventricular event also initiates a refractory period on the atrial channel – PVARP
prevent atrial sensing of a retrograde P wave

## Slide 30
PVARP
Ventricular
Channel
VSP
PAVB
Atrial Channel
AV
AB
PVAB
AS or AP
VS or VP
AP
Atrial Blanking -sense amplifier is not sensing.
prevents the pacemaker from oversensing an atrial stimulus or the depolarizing atrium
AV Delay -period from either an AS or AP event before the pacemaker paces the ventricle.
PAV-programmed 30 ms longer than the SAV
conduction of the paced atrial beat to reach the left side of the heart.
prevent sensing the ventricular beat in the atrium
prevents -oversensing.
prevent sensing any retrograde p-waves
Notes: In dual chamber devices, an AV delay is started by any paced or non-refractory atrial event and it is terminated by a sensed or paced ventricular event.
Atrial Blanking (AB) is the period of time when the sense amplifier is not sensing. This prevents the pacemaker from oversensing an atrial stimulus or the depolarizing atrium.
AV Delay is a time period from either an AS or AP event before the pacemaker paces the ventricle. The PAV is usually programmed 30 ms longer than the SAV to allow for conduction of the paced atrial beat to reach the left side of the heart.
Post Ventricular Atrial Blanking (PVAB) - The purpose of PVAB is to prevent sensing the ventricular beat in the atrium. A typical ventricular beat could be 15mV. If the atrial sensitivity is set at 0.5mV, there is a possibility for the atrial chamber to sense the R-wave. The PVAB prevents this oversensing.
Post Ventricular Atrial Refractory Period (PVARP) - The purpose of PVARP is to prevent sensing any retrograde p-waves. If the retrograde p-waves are sensed, then the patient could go into Pacemaker Mediated Tachycardia (PMT). Any atrial beat falling into PVARP is noted as AR, but not used from a timing diagram standpoint.
Atrial output also triggers a timing period on the ventricular channel, known as the ventricular blanking period

## Slide 31
TARP(total atrial refractory period)
=PVARP +AVI
limiting factor for the maximum sensed atrial rate that the pacemaker can sense and, hence, track
Notes: Total Atrial Refractory Period (TARP) - TARP is defined as the SAV+PVARP. No atrial beats can be tracked within this interval. If an atrial beat falls into this interval, then it is not tracked. When the atrial rate reaches TARP rate, the pacemaker exhibits a 2:1 block behavior.

## Slide 32
TARP
Notes: the upper rate (UR) is limited by the AVI and the PVARP. In this example, the AVI is 150 milliseconds and the PVARP is 250 milliseconds, for a TARP of 400 milliseconds (this is equal to 150ppm). As shown, after the first paced ventricular complex, a P wave occurs just after the completion of the PVARP. This P wave is sensed, initiates the AVI, and is followed by another paced ventricular complex. The subsequent P wave occurs within the PVARP and is therefore not sensed. The DDD response is to wait for the next intrinsic P wave to occur, as in this example, or for the AEI to be completed, whereupon AV sequential pacing occurs.

## Slide 33: A paced with intrinsic AV conduction- ApVs
AV interval | 
VA interval | 
LRL | 
Ap

## Slide 34: P wave tracking- AsVp
AV interval | 
VA interval | 
LRL | 
Tracking  Pacing the ventricle after an atrial intrinsic event
Notes: the pacemaker will “track” the atrium
Maintains AV Synchrony
Want to limit how fast we pace

## Slide 35: Complete inhibition- AsVs
AV interval | 
VA interval | 
LRL | 

## Slide 36: Pacemaker Crosstalk
Feature of a dual-chamber pacemaker
Detecting a paced signal in one chamber by the lead in another chamber
Inappropriate inhibition of pacing in the 2nd chamber
how to prevent crosstalk ?
BLANKING PERIOD
DECREASING ATRIAL OUTPUT
REDUCING THE VENTRICULAR SENSITIVITY
Notes: Misrepresentation of the paced signal as a cardiac depolarization signal.

## Slide 37: Atrioventricular Interval
a programmable interval initiated by a sensed or paced atrial event followed by ventricular pacing after the time interval expires.
If a ventricular event is sensed before the time interval is completed AVI will terminate and initiate VAI
PVARP
Atrial Channel
AV
AB
PVAB
AS or AP
VS or VP
Notes: he terminal portion is a relative refractory period, to assist in the detection of rapid pathologic atrial rhythms for purposes of mode switching
Even though the leading edge of the atrial pacing artifact is effectively ignored because of the blanking period, the trailing edge of the atrial pacing artifact occurring after the blanking period can occasionally be sensed on the ventricular channel. In a pacemaker-dependent patient, inhibition of ventricular output by crosstalk results in asystole. To prevent such an outcome, a safety mechanism is present

## Slide 38: Post Atrial Ventricular Blanking (PAVB)
Atrial output triggered timing period on the ventricular channel- no sensing
Prevent crosstalk.

## Slide 39
Notes: 7 Representation of atrioventricular interval (AVI). (A) AVI corresponds to a programmed value following a paced or sensed atrial beat before a ventricular pacing artifact is delivered. The post-atrial ventricular blanking period (PAVB) is found in the initial portion of the AVI, followed by the cross-talk sensing window. (B) If the ventricular-sensing amplifier senses any event during the cross-talk sensing window, ventricular safety pacing is delivered, usually 100–110 ms after the atrial event. (C) An intrinsic ventricular event [such as a premature ventricular contraction (PVC)] during the PAVB period is not sensed by the ventricular-sensing amplifier; thus ventricular pacing is delivered at the programmed AVI.

## Slide 40
Ventricular
Channel
AP
VSP
PAVB
VP
100-110 ms
PAVB: Post-Atrial Ventricular Blanking
Ventricular Safety Pacing (VSP) Window = 110ms
Atrial Channel
AS or AP
AV
AB
Notes: Post Atrial Ventricular Blanking (PAVB) - The purpose of this interval is to prevent sensing in the ventricle after an atrial paced beat - this may prevent crosstalk.
Atrial output also triggers a timing period on the ventricular channel, known as the ventricular bla
Ventricular Safety Pacing (VSP) - This is a 110 ms interval whose purpose is to prevent inhibition due to “crosstalk”.
Ventricular Blanking (VB) - This is to prevent sensing the VP or the ventricular depolarization after it has already been sensed.
Ventricular Refractory Period (VRP) - A period of time during which ventricular sensed events are ignored for timing purposes, but included in diagnostic counters. This prevents falsely sensing R-waves and T-waves, which could cause timing errors.

## Slide 41
Ventricular Safety Pacing (VSP)
Uses a brief sensing period -crosstalk sensing window
After the ventricular blanking period
Ventricular sensing during this window results in ventricular pacing at the end of the safety window
which can be recognized from the ECG by shortening of the AV interval
Notes: Any electrical ventricular event sensed during this window is regarded to be cross-talk

## Slide 43: Differential Atrioventricular Interval
Attempt to provide an intra-atrial conduction time of equal duration whether atrial contraction is paced or sensed
Paced AV interval - atrial-paced event (A) to ventricular-paced event (V)
Sensed AV interval - atrial-sensed event (P) to ventricular-paced event (V)

## Slide 44
SAV- commences only when the atrial depolarization is detected by the pacemaker.
20–60 ms after onset of P
PAV- commences immediately with the pacing artifact
SAV < PAV

## Slide 45: Dynamic AV delay
PR interval - shortened during exercise
AV delay - designed to reproduce this physiological phenomenon
Shortened sensed AV delays  increase the sensing window for detecting rapid atrial events by shortening TARP and increasing the rate of appearance of a 2:1 block.
The dynamic AV delay adjusts the AV delays linearly as the heart rate increases.
Notes: Physiologically - ,
and the same variation is applied to the sensed AV delay and the paced AV delay

## Slide 46
Possible to program 3 levels of AV delay
Low - shortening of the AV delay by 0.5 ms for a rate increase of 1 beat/minute
Medium -0.75 ms for increase of 1 beat/minute
High - 1.0 ms for increase of 1 beat/minute

## Slide 47
The shortening of the AV delay continues until the heart rate reaches the maximum sensor rate or the maximum synchronous rate or if the AV delay value reaches the value of the shortest programmed AV delay.

## Slide 48: Far field sensing
An intrinsic event is sensed in the opposite chamber it originates from
Far field R wave oversensing
Far field P wave oversensing
Notes: atrial lead picking up far-field signals from ventricular depolarization
An intrinsic event is sensed in the opposite chamber it originates from

## Slide 49: Far field R wave over sensing
Oversensing by the atrial lead
May occur with ventricular sensing or with ventricular pacing
Right atrial appendage is anatomically close to the RVOT
Notes: atrial lead picking up far-field signals from ventricular depolarization

## Slide 50
Atrial
EGM
Ventricular
EGM
Surface
ECG

## Slide 51
Consequence
- Inappropriate mode switching in a dual-chamber system
- Pacing below the programmed rate in an AAI(R) pacemaker

## Slide 52
Factors favouring
Short PVAB. FFRW falls outside of the PVAB.
Proximity of the atrial lead to the ventricle.
Unipolar sensing configuration
Programming a high atrial sensitivity level.
Interelectrode spacing of the lead (with closer spacing resulting in less FFRW oversensing)

## Slide 53
How to correct far field sensing
lengthen PVAB
Reduce atrial sensitivity
FFRW -low amplitude at implantation
Reposition the atrial lead in a more anterior or lateral position if necessary.

## Slide 54: Far field P wave oversensing
Ventricular channel senses the atrial event inappropriately
Rare

## Slide 55
Factors favouring
RV lead dislodgement to the AV junction
when an integrated bipolar lead has been positioned near the tricuspid annulus, the RV distal coil crossing the valve
RV lead unintentionally implanted in the coronary sinus
RV lead insulation defective in the atrium allowing for sensing of atrial activity

## Slide 56
In patients with sinus rhythm and complete atrio-ventricular block may inhibit ventricular pacing and lead to asystole.
Atrial flutter or atrial tachycardia  spurious shocks and possible asystole if the patient's underlying rhythm is complete atrioventricular block.
Notes: In patients with sinus rhythm and complete atrio-ventricular block, the oversensing of the P-wave may inhibit ventricular pacing and lead to asystole.
Similarly, oversensing of atrial depolarizations during atrial flutter or atrial tachycardia is likely to cause spurious shocks and possible asystole if the patient's underlying rhythm is complete atrioventricular block.

## Slide 57
How to reduce
Post-atrial paced ventricular blanking
During implantation to ensure that the distal coil lies entirely within the RV cavity

## Slide 58: Upper rate limit
Maximum tracking rate – MTR
Maximum speed limit for how fast the pacemaker can pace the heart in response to increased intrinsic atrial rate- p wave tracking
Safety feature
Don’t pace at atrial rate during AF

## Slide 59: Maximum sensor rate- MSR
Rate response mode
Used when pacemaker pace the heart in response to pacemaker sensors
Maximum rate that the pacemaker is allowed to pace at in response to what the rate response sensors are indicating

## Slide 60: Dual Chamber Timing Intervals
Ventricular
Channel
VRP
VSP*
PAVB
VB
PVARP
Atrial Channel
AV
AB
PVAB
AS or AP
VS or VP
AS or AP
VS or VP
Notes: Note: VSP window only occurs after AP

## Slide 61: Upper rate behaviour
Some timers overlap and interact with each other.
Upper rate behaviour occurs when p wave tracking at the upper rate limit or max tracking rate
Not a malfunction
Notes: With all these timers telling the pacemaker when to sense and when to pac e, some timers overlap and interact with each other
Pacemaker is bound by certain rules that it must follow making it appear as if there is something wrong with the pacemaker , when its actually a normal function.

## Slide 62
2:1 block
Pacemaker wenkebach

## Slide 63: 2:1 block
AV DELAY | 
PVARP | 
TARP | 
MTR | 
Ar
2:1 block point
= 60000 / TARP
200ms
300ms
500ms
= 60000/500
= 120 bpm
Notes: Continues till rate decreases so that every p wave lands outside the TARP

## Slide 64: Pacemaker wenkebach
AV
PVARP
TARP
MAX TRACK RATE
If atrial sense happens before MTR-
pacemaker must wait to pace the ventricle by extending the AV delay

## Slide 65
AV DELAY | 
PVARP | 
TARP | 
MTR | 
Ar
Pacemaker wenkeback ( 4:3) block

## Slide 66
Wenkeback preferred over 2: 1 block
So program accordingly
LRL
MTR
2:1 block point
1:1 conduction
wenkeback
2:1 block
As the sinus rate increases
= 60000 / TARP

## Slide 67: Pacemaker mediated tachycardia
Endless -loop tachycardia
Dual -chamber pacemakers
Characterized by atrial sensing followed by ventricular pacing at an upper tracking rate.
Notes: Endless loop and requires an outside intervention to break the loop.

## Slide 68
Reentrant circuit.
Antegrade limb - Pacemaker
Retrograde limb - AV node / accessory pathway -
Retrograde P wave produced by PVC is sensed by a pacemaker when it falls beyond the PVARP

## Slide 70
PVC
Retrograde p wave
AV
retro

## Slide 71
Requirements
-loss of AV synchrony
- V-A Retrograde conduction  AV node or an accessory pathway.
- V-A conduction > programmed PVARP

## Slide 72
Loss of AV synchrony
PVC
loss of atrial capture
oversensed / undersensed p wave
magnet application / removal
electromagnetic interference

## Slide 73
Acute termination – magnet application
Definite management-
program PVARP > retrograde conduction time

## Slide 74
AV
PVARP

## Slide 75
Programming features
adaptive PVARP in pacemakers with rate response automatically prolongs the PVARP after a PVC
Dropping a ventricular paced beat after the pacemaker had been pacing at the maximum tracking rate for a specific period of time

## Slide 76: Dynamic PVARP
PVARP-varies based on the mean atrial rate
Adaptive setting
Tracking pacing modes (DDDR, DDD, VDD)
PVARP is extended during lower heart rates in order to protect against pacemaker-mediated tachycardia
Shortened at higher rates to allow P-synchronous ventricular pacing at faster rates ; provides a higher 2:1 block rate

## Slide 77
Non-tracking pacing modes (DDI, DDIR)
PVARP is extended to promote AV synchrony at a low pacing rate
shortened to prevent atrial competitive pacing at a high pacing rate

## Slide 78: Competition and fusion
relationship between intrinsic beats and paced beats in a chamber- when both are present
competition between
intrinsic pacemaker vs implanted pacemaker
intrinsic AV conduction vs dual chamber pacing of the ventricle

## Slide 79
ventricular pacing that follows a sensed or paced atrial beat may compete with intrinsic QRS complexes – race to capture ventricle
vvi

## Slide 80: MCQs
What is the programmed Lower Rate ?
VR
VR
VR
VR
VR
VR
VR
25 mm paper speed
60 bpm
85 bpm
87 bpm
100 bpm
1000 ms
Notes: Answer: A. 60 bpm
Note that time between the two paced beats is 5 big boxes, which is equivalent to 1,000ms. (60,000ms in a minute) / (1,000ms) = 60 bpm. (A Vs is sensed as an intrinsic beat, which resets the 1,000ms timer.)
VR
VR
VR
VR
VR
VR
VR

## Slide 81
Far field R wave over sensing can be prevented by changing
VSP
PAVB
PVARP
PVAB

## Slide 82
Retrograde p wave is avoided by
VSP
PAVB
PVARP
PVAB

## Slide 83
What is the 2: 1 block point of this patient
AV= 200ms PVARP= 300ms
A-120 bpm
B- 110 bpm
C- 130 bpm
D- 100 bpm
= 60,000/ TARP= 60,000/ (AV+PVARP)

## Slide 84: Key learning points
NBG codes are used to describe the device pacing mode which dictates how the device will react to sensing
Understanding single and dual chamber timing intervals, blanking, and refractory periods is essential to understanding the operation of a pacemaker
Upper rate behavior describes the way the pacemaker responds increases in atrial rate given certain programmed AV and PVARP intervals.
In patients with heart block, upper rate timing can cause pacemaker wenckebach and/or 2:1 block.

## Slide 85
THANK YOU